## AnnexA_NewFormatTerms.csv_Final
Annex Name: Annex A
Annex Title: Medical Device Problem
Annex Version: 3.0
Annex Description: / Terms/codes for describing problems (malfunction, deterioration of function, failure) of medical devices that have occurred in pre- or post-market contexts (e.g. clinical studies, clinical evaluation or post-market surveillance)
Annex Instructions: / These terms allow capturing of the problems encountered at device(s) level through observational language without yet describing possible reasons or causes for the problems or failures observed. Annex A provides a comprehensive list of medical device problem terms and codes. It is recognized that not all jurisdictions may want to code to such detailed levels. The hierarchical structure will allow jurisdictions to choose the level of coding to use.
Annex Approval Date: 21 January 2021

Level 1 Term	Level 2 Term	Level 3 Term	Code	Definition	Non-IMDRF Code/Term	Status	Status Description	CodeHierarchy
Patient Device Interaction Problem			A01	Problem related to the interaction between the patient and the device.				A01
	Patient-Device Incompatibility		A0101	Problem associated with the interaction between the patient's physiology or anatomy and the device that affects the patient and/or the device.				A01|A0101
		Biocompatibility	A010101	Problem associated with undesirable local or systemic effects due to exposure to medical device materials or leachates from those materials by a patient who has an implant or is receiving treatment with a device made from them.				A01|A0101|A010101
		Device Appears to Trigger Rejection	A010102	The device appears to elicit undesired response in the patient to the presence of an implanted or invasive device, without inherent device failure, e.g. fibrous encapsulation, or inflammation of the tissue around the device, or extrusion of the device.				A01|A0101|A010102
		Inadequacy of Device Shape and/or Size	A010103	The physical size and/or shape of the device was inadequate with regard to the patient's anatomy.				A01|A0101|A010103
	Osseointegration Problem		A0102	Problem associated with interconnection between the bone tissue and the implanted device.				A01|A0102
		Failure to Osseointegrate	A010201	Problem associated with the failure to see direct anchorage of an implant by the formation of bony tissue around the implant without the growth of fibrous tissue at the bone-implant interface.				A01|A0102|A010201
		Loss of Osseointegration	A010202	Problem associated with weakened integration of the device at the bone-implant interface due to loss of fibrous and/or bony tissue and leading to compromised anchorage of the device. i.e. Loosening/Lysis				A01|A0102|A010202
	Loosening of Implant Not Related to Bone-Ingrowth		A0103	Problem associated with the loss of direct anchorage of an implanted device over time or due to an injury.				A01|A0103
	Migration or Expulsion of Device		A0104	Problem with an implanted or invasive device moving within the body, or being completely expelled from the body.				A01|A0104
		Expulsion	A010401	Problem with all or part of an implanted or invasive device being completely expelled from its intended location within the body.				A01|A0104|A010401
		Migration	A010402	Problem with all or part of an implanted or invasive device moving from its intended location within the body.				A01|A0104|A010402
Manufacturing, Packaging or Shipping Problem			A02	Problem associated with any deviations from the documented specifications of the device that relate to nonconformity during manufacture to the design of an item or to specified manufacturing, packaging or shipping processes (out of box problem).				A02
	Product Quality Problem		A0201	Problem associated with an inherent device characteristic that is not satisfactory as specified or delivered.				A02|A0201
		Dull, Blunt	A020101	Problem associated with a device not being as sharp as intended or expected.				A02|A0201|A020101
		Nonstandard Device	A020102	Problem associated with the device that does not meet the specifications or requirements for which it was manufactured (e.g. materials, parts, manufacturing process).				A02|A0201|A020102
	Defective Component		A0202	Problem associated with a device component having flaws of dimensional deviations greater than acceptable for the intended use.				A02|A0202
	Defective Device		A0203	Problem associated with having flaws or dimensional deviations greater than acceptable for the intended use of the device.				A02|A0203
	Device Damaged Prior to Use		A0204	Problem associated with packaging or shipping damage prior to the use of the device.				A02|A0204
	Packaging Problem		A0205	Problem associated with the materials used to construct the cover or outer wrapping of the device.				A02|A0205
		Difficult to Open or Remove Packaging Material	A020501	Problem associated with difficulty for users to operate the device, specifically as it relates to the opening or removal of the outer wrapping.				A02|A0205|A020501
		Incomplete or Missing Packaging	A020502	Problem associated with the nonconformance to the device specifications due to incomplete or missing packaging that may compromise the device operation as intended.				A02|A0205|A020502
		Unsealed Device Packaging	A020503	Problem associated with the loss of packaging seal.				A02|A0205|A020503
		Tear, Rip or Hole in Device Packaging	A020504	Problem associated with packaging damage (tear, rip or hole) prior to the use of the device.				A02|A0205|A020504
	Device Misassembled During Manufacturing /Shipping		A0206	A device found incorrectly assembled when delivered to the user facility.				A02|A0206
		Component Misassembled	A020601	A device found to have one or more components incorrectly assembled when delivered to the user facility.				A02|A0206|A020601
		Component Missing	A020602	A device component(s) found to be missing when delivered to the user facility.				A02|A0206|A020602
	Shipping Damage or Problem		A0207	Problem associated with shipping damage or problem prior to the use of the device.				A02|A0207
		Delivered as Unsterile Product	A020701	Problem associated with a device being received in such a manner to indicate that its sterility has been compromised (e.g. sterile packaging breached, visible contaminate present)				A02|A0207|A020701
Chemical Problem			A03	Problem associated with any from the documented specifications of the device that relate to any chemical characterization, i.e., element, compound, or mixture.				A03
	Device Emits Odor		A0301	Problem associated with an unexpected or inappropriate smell released by the device.				A03|A0301
	Device Ingredient or Reagent Problem		A0302	Problem associated with any deviations from the documented specifications of the device that relate to any ingredient or reagent characterization.				A03|A0302
		Clumping in Device or Device Ingredient	A030201	Problem associated with the aggregation of particles into irregular masses.				A03|A0302|A030201
		Coagulation in Device or Device Ingredient	A030202	Problem associated with the undesired characterization of congealing, solidifying, thickening, curdling.				A03|A0302|A030202
		Precipitate in Device or Device Ingredient	A030203	Problem associated with the separation of solid particles from a liquid as the result of a chemical or physical change.				A03|A0302|A030203
		Cross Reactivity	A030204	Problem associated with the degree to which an antibody or antigen participates in cross reactions.				A03|A0302|A030204
		Particulates	A030205	Substances that consist of separate particles that are introduced by the device during use.				A03|A0302|A030205
		High pH	A030206	pH higher than expected and/or anticipated.				A03|A0302|A030206
		Low pH	A030207	pH lower than expected and/or anticipated.				A03|A0302|A030207
		Unexpected Color	A030208	Color of product is different from that expected.		New	Term was added on 18 March 2020 as an additional level 3 term under "Device Ingredient or Reagent Problem".	A03|A0302|A030208
	Improper Chemical Reaction		A0303	Problem associated with an unexpected or incomplete chemical reaction or effect.				A03|A0303
Material Integrity Problem			A04	Problem associated with any deviations from the documented specifications of the device that relate to the limited durability of all material used to construct device.				A04
	Break		A0401	Problem associated with undesired damage or breakage of those materials used in the device construction.				A04|A0401
		Fracture	A040101	Problem associated with a partial or full-thickness crack in the device materials.				A04|A0401|A040101
		Loss of or Failure to Bond	A040102	Problem associated with lack or loss of adherence between materials intended to be joined together by an adhesive.				A04|A0401|A040102
		Material Fragmentation	A040103	Problem associated with small pieces of the device breaking off unexpectedly.				A04|A0401|A040103
		Solder Joint Fracture	A040104	Problem associated with undesired damage or breakage in a solder joint of materials used in the device construction.				A04|A0401|A040104
	Burst Container or Vessel		A0402	Problem associated with the pressure inside a vessel or container rising to such a degree that the container ruptures.				A04|A0402
	Explosion		A0403	Problem associated with the violent bursting due to the sudden expansion of air, gas or fluid.				A04|A0403
	Crack		A0404	Problem associated with an undesired partial separation and/or a visible opening along the length or width in the materials that are used in the device construction.				A04|A0404
	Degraded		A0405	Problem associated with a undesired change in the chemical structure, physical properties, or appearance in the materials that are used in the device construction.				A04|A0405
		Calcified	A040501	Problem associated with buildup of calcium salts on the device.				A04|A0405|A040501
		Corroded	A040502	Problem associated with the chemical or electrochemical reaction between materials, usually a metal and its environment that produces a deterioration of the metal and its properties.				A04|A0405|A040502
		Material Erosion	A040503	Problem associated with a progressive loss of a material from a solid surface.				A04|A0405|A040503
		Pitted	A040504	Problem associated with the corrosion of a material's surface, confined to a point or small area that takes the form of cavities.				A04|A0405|A040504
		Flaked	A040505	Problem associated with the detachment of small pieces of the coating film of a material.				A04|A0405|A040505
		Peeled/Delaminated	A040506	Peeling or delamination of composite materials, including coatings, that occurs when layers are separated as a result of stress or impact and resulting in loss of mechanical toughness.				A04|A0405|A040506
		Naturally Worn	A040507	Problem associated with material damage to a surface, usually involving progressive loss or displacement of material, due to relative motion between that surface and a contacting substance or substances.				A04|A0405|A040507
		Unraveled Material	A040508	Problem due to the undesired unravelling of material (e.g. disentangled, unwound etc.).				A04|A0405|A040508
	Material Deformation		A0406	Problem associated with an undesired material change in shape or property caused by external forces.				A04|A0406
		Deformation Due to Compressive Stress	A040601	Problem associated with an undesired bulge, bend, bow, kink, or wavy condition observed in the device material resulting from compressive stresses.				A04|A0406|A040601
		Dent in Material	A040602	Problem associated with a undesired change in shape, characterized by the presence of a slight hollow (dent) in the device surface.				A04|A0406|A040602
		Failure to Fold	A040603	Problem associated with an undesired material change in physical property, characterized by failure to fold.				A04|A0406|A040603
		Failure to Unfold or Unwrap	A040604	Problem associated with the comprising materials' deformation in that device fails to open its wrapping or open/extend in a certain manner i.e. balloon or lens.				A04|A0406|A040604
		Material Frayed	A040605	Problem associated with the comprising materials having damaged edges.				A04|A0406|A040605
		Material Invagination	A040606	Problem associated with an undesired material change in shape, characterized by the infolding of one part within another part of a structure.				A04|A0406|A040606
		Material Too Rigid or Stiff	A040607	Problem associated with an undesired material change in physical property, characterized by rigidity (it resists deformation in response to an applied force).				A04|A0406|A040607
		Material Too Soft/Flexible	A040608	Problem associated with any device material that results in the material's inability to maintain the desired shape or support function.				A04|A0406|A040608
		Material Twisted/Bent	A040609	Problem associated with deformations that lead to twisting or bending of the device.				A04|A0406|A040609
		Melted	A040610	Problem associated with a solid device being transformed into a molten or liquid state.				A04|A0406|A040610
		Stretched	A040611	Problem associated with an increase or elongation in a materials' dimension.				A04|A0406|A040611
		Thickening of Material	A040612	Device is thickened due to deformation. Note: Not to be used for thickening of liquid.		New	Term was added on 18 March 2020 as an additional level 3 term under "Material Deformation".	A04|A0406|A040612
	Material Discolored		A0407	Problem associated with an undesired streak, pattern and/or a noticeable change in color from the rest of the materials used in the device construction.				A04|A0407
	Material Disintegration		A0408	Problem associated with material breaking into small particles.				A04|A0408
	Material Opacification		A0409	Problem associated with an undesirable opaqueness or cloudiness.				A04|A0409
	Material Perforation		A0410	Material constituting device is perforated possibly compromising the device's intended purpose.				A04|A0410
		Material Puncture/Hole	A041001	Device material(s) punctured leading to undesired holes/openings.				A04|A0410|A041001
	Material Protrusion/Extrusion		A0411	Problem associated with undesired physical appearance of device material, specifically when material extends beyond or above device surface.				A04|A0411
	Material Rupture		A0412	Problem associated with perforations that lead to bursting of the device.				A04|A0412
	Material Separation		A0413	Problem associated with an undesired disassociation or breaking apart of the device.				A04|A0413
	Material Split, Cut or Torn		A0414	Problem associated with materials consisting the device are split, cut or torn due to external forces (e.g. wrenching or laceration) or internal forces (e.g. exceeding the tensile stress limits belonging to the materials used in the device construction).				A04|A0414
	Scratched Material		A0415	Problem associated with an undesirable shallow cut or narrow groove in the surface of the device materials.				A04|A0415
Mechanical Problem			A05	Problems associated with mechanical actions or defects, including moving parts or subassemblies, etc.				A05
	Detachment of Device or Device Component		A0501	Problem associated with the separation of the device from its physical construct, integrity, or chassis.				A05|A0501
	Device Damaged by Another Device		A0502	Problem associated with one device causing harm to another device.				A05|A0502
	Ejection Problem		A0503	Problems associated with the inability or unexpected ejection or discharge of the device from its physical location.		Modified (editorial)	Editorial change to the definition on 18 March 2020.	A05|A0503
		Failure to Eject	A050301	Problem associated with the inability of the device to be ejected or discharged from its physical location.		Modified (editorial)	Editorial change to the definition on 18 March 2020.	A05|A0503|A050301
		Unintended Ejection	A050302	Problem associated with unexpected ejection of the device from its physical location. This includes but is not limited to devices such as clip appliers, film cartridge, and staples.		Modified (editorial)	Editorial change to the definition on 18 March 2020.	A05|A0503|A050302
	Leak/Splash		A0504	Problem associated with the escape of a liquid, gas or radiation from the vessel or container in which it is housed.		Modified (technical)	Technical change to the definition on 18 March 2020 to include escape of radiation.	A05|A0504
		Fluid Leak	A050401	Escape (Release, Discharge) of fluid through an unintended location - often accompanied by a loss of pressure and/or output.				A05|A0504|A050401
		Gas Leak	A050402	Problem associated with the unintended escape of a gas from the container in which it is housed.				A05|A0504|A050402
		Gel Leak	A050403	Escape (Release, Discharge) of gel through an unintended location - as in leakage of ultrasound gel. Escape or release of gel from containment structures - as in gel filled implant leak.				A05|A0504|A050403
		Radiation Leak	A050404	Escape of radiation (energy in the form of waves or subatomic particles, especially those that cause ionization) through containment structures, leading to unintended exposure.				A05|A0504|A050404
		Perivalvular Leak	A050405	Problem associated with the escape of blood around a heart valve, particularly around its leaflets.				A05|A0504|A050405
	Firing Problem		A0505	Problem associated with the device not discharging as intended.				A05|A0505
		Failure to Fire	A050501	Problem associated with failure of the device to discharge its load (e.g. surgical stapler failed to partially or completely deploy its staples).				A05|A0505|A050501
		Misfire	A050502	Problem associated with a therapy or algorithm not being delivered or executed at the expected time.				A05|A0505|A050502
	Mechanical Jam		A0506	The motion of the device is prevented or restricted.				A05|A0506
	Mechanics Altered		A0507	Problem associated with a device mechanical functioning of machinery, moving parts or tools of device being changed or modified.				A05|A0507
		Failure to Align	A050701	Problem associated with a circuit, equipment, or system whereby its functions fail to be properly synchronized or its relative positions properly oriented.				A05|A0507|A050701
		Failure to Cut	A050702	Inability of the device to make an incision, pierce or open as intended.				A05|A0507|A050702
		Failure to Cycle	A050703	Problem associated with the device failing to complete a series of processes or events.				A05|A0507|A050703
		Failure to Form Staple	A050704	Problem associated with the device failing to connect tissue with a stapling device due to the staples not forming correctly.				A05|A0507|A050704
	Noise, Audible		A0508	Problem associated with any unintended sound which emanates from the device (for example, squeaking from two parts rubbing together or buzzing sounds from electrical components).				A05|A0508
	Physical Resistance/Sticking		A0509	Problem associated with the lack of movement in the device due parts sticking or seizing.				A05|A0509
	Retraction Problem		A0510	Problem associated with drawing back the device to an intended location.				A05|A0510
	Structural Problem		A0511	Problem associated with the basic physical construction or physical make up of the device.				A05|A0511
		Collapse	A051101	Problem associated with the buckling or crushing of material from external forces.				A05|A0511|A051101
		Sharp Edges	A051102	The device has undesirable sharp edges which can cause harm or damage.				A05|A0511|A051102
		Difficult to Fold, Unfold or Collapse	A051103	Problem associated with the use of the device in terms of user experiencing difficulty to close or to spread out/extend length of the device, even if the operation is being performed according to labeled instructions for use.		Modified (technical)	Technical change to the term on 18 March 2020 to include difficulty in collapsing.	A05|A0511|A051103
		Difficult to Open or Close	A051104	Problem associated with the use of the device in terms of user experiencing difficulty opening and closing the device, even if the operation is being performed according to labeled instructions for use.				A05|A0511|A051104
		Incomplete Coaptation	A051105	Problem associated with the heart valve leaflet not closing properly.				A05|A0511|A051105
	Unintended Movement		A0512	Problem associated with an undesired movement of the device, which may be related to the device malfunction, misdiagnosis, or mistreatment.				A05|A0512
		Device Dislodged or Dislocated	A051201	Problems associated with the device not remaining in an expected location.				A05|A0512|A051201
		Device Tipped Over	A051202	Problem associated with the inability of the device to stay in an upright position.				A05|A0512|A051202
		Device Fell	A051203	Problem associated with the device or a component unexpectedly being dropped or moving down from an intended place.				A05|A0512|A051203
		Device Slipped	A051204	Problem associated with the device moving or sliding from the intended position.				A05|A0512|A051204
		Unintended Collision	A051205	Problem associated with the device impacting with another object.				A05|A0512|A051205
		Unintended System Motion	A051206	Problem associated with any motion of the system or components that was not initiated by the user.				A05|A0512|A051206
		Unstable	A051207	Problem associated with the mechanical stability of the device.				A05|A0512|A051207
		Vibration	A051208	Problem associated with the undesirable mechanical oscillation.				A05|A0512|A051208
	Inadequate Lubrication		A0513	The device shows inadequate lubrication.		New	Term was added on 18 March 2020 as an additional level 2 term under "Mechanical Problem."	A05|A0513
Optical Problem			A06	Problem associated with transmission of visible light affecting the quality of the image transmitted or otherwise affecting the intended application of the visible light path.				A06
	Misfocusing		A0601	The problem relates to the poor focusing of the object or the focus is on the wrong object or in the wrong area.				A06|A0601
	Optical Decentration		A0602	Problem associated with being off-center of optical lenses.				A06|A0602
	Optical Discoloration		A0603	Problem associated with an undesired change of color.				A06|A0603
	Optical Distortion		A0604	Problem associated with an optical defect in an image-forming system whereby the image is not the shape of an ideal image of the object.				A06|A0604
	Optical Obstruction		A0605	Problem associated with the blocking of optical devices, e.g. visual pathways.				A06|A0605
Electrical /Electronic Property Problem			A07	Problem associated with a failure of the electrical circuitry of the device.				A07
	Capturing Problem		A0701	Problem associated with the inability of the device to achieve successful depolarization and contraction of a cardiac chamber caused by a pacemaker output pulse.				A07|A0701
		Failure to Capture	A070101	Problem associated with the failure to achieve effective and consistent depolarization of the heart resulting from the electrical stimulus of the pacemaker.				A07|A0701|A070101
		High Capture Threshold	A070102	Problem with the amount of output energy needed to cause cardiac depolarization being higher than expected/desired.				A07|A0701|A070102
		Intermittent Capture	A070103	Problem associated with the ineffective and inconsistent depolarization of the heart.				A07|A0701|A070103
		Unstable Capture Threshold	A070104	Problem with the amount of output energy needed to cause cardiac depolarization being unstable.				A07|A0701|A070104
	Continuous Firing		A0702	Problem associated with the excessive production of electrical impulses over a period.				A07|A0702
	Arcing		A0703	Problem associated with electrical current flowing through a gap between two conductive surfaces, typically resulting in a visible flash of light.				A07|A0703
		Arcing at Paddles	A070301	Problem associated with electrical current flowing through a gap between paddles (conductive surfaces), typically resulting in a visible flash of light.				A07|A0703|A070301
		Arcing of Electrodes	A070302	Problem associated with electrical current flowing through a gap between electrodes (conductive surfaces), typically resulting in a visible flash of light.				A07|A0703|A070302
	Sparking		A0704	Problem associated with a flash of light related to an electrical discharge into a normally non conductive medium, such as air. Not associated with a discharge between two conductive surfaces.				A07|A0704
	Battery Problem		A0705	Problem associated with the internal power of the device (e.g. battery, transformer, fuel cell or other power sources).				A07|A0705
		Battery Problem: High Impedance	A070501	Problem related to increased battery internal impedance.				A07|A0705|A070501
		Battery Problem: Low Impedance	A070502	Problem related to decreased battery internal impedance.				A07|A0705|A070502
		Failure to Run on Battery	A070503	Problem associated with the device failing to operate when not connected to a fixed power source.				A07|A0705|A070503
		Premature Discharge of Battery	A070504	Battery discharging earlier than expected.				A07|A0705|A070504
	Charging Problem		A0706	Problem associated with the inability of the device to successfully charge an electrical source.				A07|A0706
		Aborted Charge	A070601	Problem associated with the premature ending of the charging process (e.g. of a battery or other charge storage device).				A07|A0706|A070601
		Delayed Charge Time	A070602	Problem associated with an unexpected amount of time required to charge the device (e.g. a delay in starting charging or a longer than expected charge time).				A07|A0706|A070602
		Failure to Charge	A070603	Problem associated with inability to initiate the appropriate charging process (e.g. of a battery or other charge storage device)				A07|A0706|A070603
	Failure to Discharge		A0707	Problem associated with the failure of a battery or other charge storage device to appropriately discharge as intended. Does not apply to defibrillation.				A07|A0707
	Power Problem		A0708	Problem associated with the energy to operate the device.				A07|A0708
		Complete Loss of Power	A070801	Problem associated with the lack of power to run the device.				A07|A0708|A070801
		Intermittent Loss of Power	A070802	Problem associated with an intermittent disruption to the power to run the device.				A07|A0708|A070802
		Failure to Power Up	A070803	Problem associated with the inability of the device to turn on related to energy delivered to the device.				A07|A0708|A070803
		Unintended Power Up	A070804	Problem associated with the device turning on when not intended.				A07|A0708|A070804
	Device Sensing Problem		A0709	Problem associated with the device feature that are designed to respond to a physical stimulus (temperature, illumination, motion, cardiac rhythms) and that do not transmit a resulting signal for interpretation or measurement.				A07|A0709
		Decreased Sensitivity	A070901	Problem with the device being less sensitive to an input than intended or expected.				A07|A0709|A070901
		Increased Sensitivity	A070902	Problem with the device being more sensitive to an input than intended or expected.				A07|A0709|A070902
		Failure to Analyze Signal	A070903	Problem with the device not analyzing a signal.				A07|A0709|A070903
		Failure to Select Signal	A070904	Problem associated with the failure of the device to select the appropriate input signal.				A07|A0709|A070904
		High Sensing Threshold	A070905	Problem associated with the amount of input required by the device to detect a signal being higher than expected/desired.				A07|A0709|A070905
		Low Sensing Threshold	A070906	Problem associated with the amount of an input required by the device to detect a signal being lower than expected/desired.				A07|A0709|A070906
		Loss of Threshold	A070907	Problem associated with the loss of the minimum amount of energy, voltage, or current needed to consistently stimulate the heart muscle.				A07|A0709|A070907
		Failure to Sense	A070908	Problem associated with the failure of the device designed to respond to a physical stimulus (as temperature, illumination, motion) to transmit a resulting signal for interpretation or measurement.				A07|A0709|A070908
		Over-Sensing	A070909	Problem related to failure of the device to properly filter cardiac signals resulting in inappropriate device response.				A07|A0709|A070909
		Under-Sensing	A070910	Problem related to failure of the device to properly detect intrinsic cardiac activity and respond appropriately.				A07|A0709|A070910
		Sensing Intermittently	A070911	Problem with the device receiving an incoming signal on an intermittent basis when expected to be continuous.				A07|A0709|A070911
		Incorrect Interpretation of Signal	A070912	Problem with the device inappropriately analyzing a signal.				A07|A0709|A070912
	Failure to Conduct		A0710	Problem associated with the inability of the device to allow a current of electricity to pass or to conduct electricity continuously along an electrical path.				A07|A0710
	Interrogation Problem		A0711	Problems associated with the device's ability to respond to signals from a system designed to interrogate its status.				A07|A0711
		Difficult to Interrogate	A071101	Problem associated with difficulty of a transponder system to trigger a response.				A07|A0711|A071101
		Failure to Interrogate	A071102	Problem associated with the device failure to appropriately respond to signals from a system designed to interrogate its status.				A07|A0711|A071102
	Pacing Problem		A0712	Problem associated with the inability of the device to generate a therapeutic simulated heart beat via electrical impulses.				A07|A0712
		Failure to Convert Rhythm	A071201	Failure of the device therapy or set of therapies to terminate the harmful cardiac rhythm that the therapy is meant to terminate.				A07|A0712|A071201
		Inaccurate Synchronization	A071202	Problem associated with an error due to imperfect timing of two operations, e.g. signal transmission time.				A07|A0712|A071202
		Inappropriate Waveform	A071203	Failure of the device to generate a correctly-shaped pacing output, e.g., a waveform that is too wide.				A07|A0712|A071203
		No Pacing	A071204	Problem associated with the device ceasing to deliver paces.				A07|A0712|A071204
		Pacemaker Found in Back-Up Mode	A071205	A device with a pacing function found in back-up Mode. This may be an appropriate fail-safe action (e.g. end of battery life), or be caused by device malfunction or due to operator error.				A07|A0712|A071205
		Pacing Asynchronously	A071206	Problem associated with a pacing transmission process such that between any two significant instants in the same group, there is always an integral number of unit intervals. Between two significant instants located in different groups, there are not always an integral number of unit intervals.				A07|A0712|A071206
		Pacing Inadequately	A071207	Pacing voltage or pulse width is less than desired.				A07|A0712|A071207
		Pacing Intermittently	A071208	Problem associated with the failure of pacing device for a limited period of time, following which the item recovers its ability to perform its required function without being subjected to any external corrective action. Note: such as failure is often recurrent.				A07|A0712|A071208
		Pocket Stimulation	A071209	Problem associated with a pocket of skin in which the pulse generator is housed.				A07|A0712|A071209
	Defibrillation/Stimulation Problem		A0713	Problem associated with the delivery of electrical shock/stimulation by a device intended to deliver shocks/stimulation.		Modified (technical)	Technical change to the term and definition on 18 March 2020 to include stimulation and to differentiate this term with "Unintended Electrical Shock" (A0714).	A07|A0713
		Failure to Deliver Shock/Stimulation	A071301	Problem associated with the failure of the device to deliver electrical energy.		Modified (technical)	Technical change to the term on 18 March 2020 to include stimulation.	A07|A0713|A071301
		Inappropriate/Inadequate Shock/Stimulation	A071302	A shock was delivered with incorrect energy level or at incorrect timing.		Modified (technical)	Technical change to the term and definition on 18 March 2020 to include inadequate stimulation.	A07|A0713|A071302
		Intermittent Shock/Stimulation	A071303	Problem associated with the failure to deliver electrical energy for a limited period of time, following which the item recovers its ability to perform its required function without being subjected to any external corrective action. Note: such a failure is often recurrent.		Modified (technical)	Technical change to the term and definition on 18 March 2020 to include stimulation.	A07|A0713|A071303
	Unintended Electrical Shock		A0714	The device delivers unintended electrical shock. Use this term for devices that are not intended to deliver shocks/stimulation.		Modified (technical)	Technical change to the definition on 18 March 2020 to differentiate this term with "Defibrillation/Stimulation Problem" (A0713) and lower level terms.	A07|A0714
	Grounding Malfunction		A0715	Problem associated with the inability to connect conductors of an electronic system for the purpose of controlling or impeding ground currents and voltages.				A07|A0715
	Electrical Overstress		A0716	Problem associated with an electrical activity that exceeded the specified threshold limit of the internal integrated circuitry.				A07|A0716
	Electro-Static Discharge		A0717	Problem associated with the discharge of electricity between two bodies previously electrically charged.				A07|A0717
	Failure to Shut Off		A0718	Problem associated with the device not powering off when a shut down was requested.				A07|A0718
		Device Remains Activated	A071801	Problem associated with the device continuing to be in an active state after deactivation was requested.				A07|A0718|A071801
	Unexpected Shutdown		A0719	Problem associated with the device unexpectedly powering down.				A07|A0719
	Electromagnetic Compatibility Problem		A0720	Problem associated with the ability of a system to function in its electromagnetic environment without introducing intolerable disturbances to anything in its environment.				A07|A0720
		Electromagnetic Interference	A072001	Problem associated with a measure of electromagnetic radiation from equipment.				A07|A0720|A072001
		Radiofrequency Interference (RFI)	A072002	Problem associated with the degradation of the reception of a wanted signal caused by RF disturbance.				A07|A0720|A072002
	Circuit Failure		A0721	Problem associated with a failure of the internal network paths or electrical circuitry (i.e. electrical components, circuit boards, wiring)				A07|A0721
		Capacitative Coupling	A072101	Problem associated with the transfer of energy within an electrical network by means of the capacitance between circuit nodes. It occurs when energy is coupled from one circuit to another through an electric field.				A07|A0721|A072101
		Electrical Shorting	A072102	Problem associated with an electric current travelling along an accidental path (unintended path) in a circuit.				A07|A0721|A072102
		Intermittent Continuity	A072103	Problem associated with intermittent faults in electrical/electronic interconnections.				A07|A0721|A072103
	Impedance Problem		A0722	Problem associated with electrical impedance levels between device and patient connections.				A07|A0722
		High impedance	A072201	Problem associated with higher than intended electrical impedance levels between device and patient connections.				A07|A0722|A072201
		Low impedance	A072202	Problem associated with lower than intended electrical impedance levels between device and patient connections.				A07|A0722|A072202
Calibration Problem			A08	Problem associated with the operation of the device, related to its accuracy, and associated with the calibration of the device.				A08
	Failure to Calibrate		A0801	Problem associated with the failure of the device to perform a self-calibration procedure or process designed to assure the accuracy and proper performance of the device.				A08|A0801
	Failure to Recalibrate		A0802	Problem associated with the failure of the device which is unable to regain a standard level of accuracy when performing a calibration procedure or process designed to assure the accuracy and proper performance of the device.				A08|A0802
	Imprecision		A0803	Problem associated with the device providing imprecise measurements when compared to a reference standard.				A08|A0803
	Overcorrection		A0804	Problem associated with an adjustment that surpasses a set of criteria.				A08|A0804
	Undercorrection		A0805	Problem associated with an adjustment that falls below a set of criteria.		New	Term was added on 18 March 2020 as an additional level 2 term under "Calibration Problem".	A08|A0805
Output Problem			A09	Problem associated with any deviation from the documented specifications of the device that relate to the end result, data, or test results provided by the device.				A09
	Audible Prompt/Feedback Problem		A0901	Problem with any deviation from the documented specifications of the device that relate to audible feedback. e.g. voice prompts or beeps, but not safety related alarms which are covered under "Protective Measures Problem".		Modified (editorial)	Editorial change to the term on 18 March 2020.	A09|A0901
		Inappropriate Audible Prompt/Feedback	A090101	Problem with audible messages which do not guide a device user to the correct action.				A09|A0901|A090101
		Inaudible or Unclear Audible Prompt/Feedback	A090102	Problem associated with audible prompts which cannot be heard clearly.				A09|A0901|A090102
		No Audible Prompt/Feedback	A090103	Problem associated with the device ceasing to provide audible prompts.				A09|A0901|A090103
	Display or Visual Feedback Problem		A0902	Problem with any deviation from the documented specifications of the device that relate to visual feedback. e.g. the display of information, images on a screen, or output from the device.				A09|A0902
		Device Displays Incorrect Message	A090201	Problem associated with providing incorrect display information.				A09|A0902|A090201
		Display Difficult to Read	A090202	Problem associated with legibility of the display, compromising for instance the reading/interpretation of patient parameters or test results. Legibility problems can be due to color, size of font, display screen contrast or other factors.				A09|A0902|A090202
		Erratic or Intermittent Display	A090203	A device does not consistently display the same message, result, reading, or image. e.g. the display might flicker, switch between readings or messages, or go completely blank for brief periods of time.				A09|A0902|A090203
		Image Display Error/Artifact	A090204	Problem with image display leading to corrupted images or readouts/measurement indications.				A09|A0902|A090204
		Image Orientation Incorrect	A090205	Problem associated with an incorrect image orientation on the device display.				A09|A0902|A090205
		No Display/Image	A090206	Problem associated with the absence of display or image.				A09|A0902|A090206
		No Visual Prompts/Feedback	A090207	Problem associated with the device ceasing to provide visual feedback.				A09|A0902|A090207
		Poor Quality Image	A090208	Inadequate quality of an image or any visual representation displayed by the device, or output from the device.				A09|A0902|A090208
		Visual Prompts will not Clear	A090209	Problem with visual messages which continue to be displayed on/by the device after the appropriate action has been taken.				A09|A0902|A090209
	Tactile Prompts/Feedback		A0903	Problem with any deviation from the documented specifications of the device that relate to tactile feedback. e.g device vibrational prompt.				A09|A0903
		Inappropriate Tactile Prompt/Feedback	A090301	Problem with tactile feedback which does not guide a device user to the correct action.				A09|A0903|A090301
		No Tactile Prompts/Feedback	A090302	Problem associated with the device ceasing to provide tactile feedback.				A09|A0903|A090302
	Energy Output Problem		A0904	Problem with the device's intended output of energy.				A09|A0904
		Energy Spectrum Incorrect	A090401	Problem associated with the energy output from the device not being in the expected part of the spectrum.				A09|A0904|A090401
		Failure to Deliver Energy	A090402	Problem associated with the failure of the device to deliver any energy.				A09|A0904|A090402
		Intermittent Energy Output	A090403	Problem associated with the energy output from the device being inconsistent over time.				A09|A0904|A090403
		Output above Specifications	A090404	Device output is exceeding the documented specifications of the device.				A09|A0904|A090404
		Output below Specifications	A090405	Device output is below the documented specifications of the device.				A09|A0904|A090405
		Therapeutic or Diagnostic Output Failure	A090406	Problem associated with the failure of the device to deliver the output required for treatment or identification of a disease.				A09|A0904|A090406
		Therapy Delivered to Incorrect Body Area	A090407	Problem associated with the device causing unintended therapeutic action to an area of the body other than the intended area.				A09|A0904|A090407
	Radiation Output Problem		A0905	Problem with the device's intended output of radiation.				A09|A0905
		Radiation Output Failure	A090501	Problem associated with the absence of radiation output from radiological or diagnostic devices.				A09|A0905|A090501
		Radiation Overexposure	A090502	Problem associated with excessive radiation emitted from radiological or diagnostic devices.				A09|A0905|A090502
		Radiation Underexposure	A090503	Problem associated with too little radiation emitted from radiological or diagnostic devices.				A09|A0905|A090503
		Unexpected/Unintended Radiation Output	A090504	Device-emitted radiation when it was not supposed to. This applies to devices which are intended to emit radiation, and the radiation being emitted from the correct part of the device, but at an incorrect time. Use "radiation leak" if the device emits radiation which should never have been emitted, or from a location from which it should never be emitted.				A09|A0905|A090504
	Gas Output Problem		A0906	Problem associated with gas output.				A09|A0906
	No Device Output		A0907	Problem associated with no measurement outcome, value or data obtained from the device.				A09|A0907
	Incorrect, Inadequate or Imprecise Result or Readings		A0908	Problem associated with a nonconforming end result, data, or test results provided by the device to its performance specifications.		Modified (editorial)	Editorial change to the term on 16 Feburary 2021.	A09|A0908
		Signal Artifact/Noise	A090801	Problem associated with impurities or interference in a signal (e.g. ECG artifact).		Modified (technical)	Technical change to the term on 18 March 2020 to include signal noise.	A09|A0908|A090801
		Failure to Obtain Sample	A090802	The device does not collect or transfer the sample.				A09|A0908|A090802
		False Negative Result	A090803	Problem associated with the device incorrectly reporting that something has not been detected and may mislead the operator into not taking certain actions when action should be taken.				A09|A0908|A090803
		False Positive Result	A090804	Problem associated with the device incorrectly reporting that something has been detected and may mislead the operator to take certain actions.				A09|A0908|A090804
		Incorrect Measurement	A090805	Measurement obtained from or provided by the device is obviously incorrect.				A09|A0908|A090805
		Non Reproducible Results	A090806	Device results cannot be reliably reproduced.				A09|A0908|A090806
		High Readings	A090807	Reading provided by the device is too high or higher than expected.				A09|A0908|A090807
		Low Readings	A090808	Reading provided by the device is too low or lower than expected.				A09|A0908|A090808
		High Test Results	A090809	Test results provided by the device are too high or higher than expected.				A09|A0908|A090809
		Low Test Results	A090810	Test results provided by the device are too low or lower than expected.				A09|A0908|A090810
		Unable to Obtain Readings	A090811	The device does not provide or display a valid reading.				A09|A0908|A090811
		Missing Test Results	A090812	Problem associated with the results of a test or measurement not appearing.				A09|A0908|A090812
		Erratic Results	A090813	Reports of erroneous/discrepant results which combine high/low and/or positive/negative results. This term is not to be selected where reports indicate consistently high or low or false positive or false negative results.		New	Term was added on 18 March 2020 as an additional level 3 term to "Incorrect, Inadequate or Imprecise Results or Readings".	A09|A0908|A090813
	Unexpected Therapeutic Results		A0909	Problem associated with the use of the device for therapeutic purposes.				A09|A0909
Temperature Problem			A10	Problem associated with the device producing unintended temperatures.				A10
	Excessive Cooling		A1001	Problem associated with the device producing temperatures that are lower than specified.				A10|A1001
	Excessive Heating		A1002	Problem associated with the device which have a warming or heating function, producing excessive heat.				A10|A1002
	Insufficient Cooling		A1003	Problem associated with the device insufficiently cooled in device active (working) or/and non-active (nonworking) state.				A10|A1003
	Insufficient Heating		A1004	Problem associated with the device or its components producing temperatures that are not as high as what is specified.				A10|A1004
	Overheating of Device		A1005	Problem associated with the device producing high temperatures, such that its operation is compromised or harm is caused (e.g. overheating that produces melting of components or automatic shutdown).				A10|A1005
	Thermal Decomposition of Device		A1006	Problems associated with a discoloration or destruction as a result of thermal decomposition of the device.				A10|A1006
	Fire		A1007	Problem associated with the combustion of the device with a steady flame.				A10|A1007
	Flare or Flash		A1008	Problem associated with device-related burn with an unsteady flame.				A10|A1008
	Smoking		A1009	Problem associated with a cloud of vapor or gas generated from the device, generally associated after a fire or a burn.				A10|A1009
Computer Software Problem			A11	Problem associated with written programs, codes, and/or software system that affects device performance or communication with another device.				A11
	Application Network Problem		A1101	Problem associated with the deviations from documented system specifications that affects overall system performance and/or the performance of an individual device connected to that system.				A11|A1101
	Application Program Problem		A1102	Problem associated with the requirement for software to fulfill its function within an intended use or application.				A11|A1102
		Application Program Freezes, Becomes Nonfunctional	A110201	Problem associated with freezing and becoming nonfunctional of an application program.				A11|A1102|A110201
		Application Program Problem: Dose Calculation Error	A110202	Problem associated with the written program code or application software used by the device to calculate specific measurements or quantities managed by the device.				A11|A1102|A110202
		Application Program Problem: Medication Error	A110203	Event in which the device software results in errors of medication preparation or administration.				A11|A1102|A110203
		Application Program Problem: Parameter Calculation Error	A110204	Problem associated with the written program code or application software used by the device to calculate parameters other than those related to dose or power.				A11|A1102|A110204
		Application Program Problem: Power Calculation Error	A110205	Problem associated with the written program code or application software used by the device for calculations related to device power.				A11|A1102|A110205
		Application Program Version or Upgrade Problem	A110206	Problem associated with installing updates to a software system that affects the device performance or communication with another device.				A11|A1102|A110206
		Problem with Software Installation	A110207	Problem associated with installing the device software in a manner that allows full functioning of the device. Source of installation could be manufacturer or user.				A11|A1102|A110207
		Unintended Application Program Shut Down	A110208	Problem associated with an unintended shut down by malfunction of the application program.				A11|A1102|A110208
	Program or Algorithm Execution Problem		A1103	Problem associated with execution problems relating to program or algorithm.				A11|A1103
		Delayed Program or Algorithm Execution	A110301	Problem associated with delayed execution relating to program or algorithm.				A11|A1103|A110301
		Intermittent Program or Algorithm Execution	A110302	Problem associated with intermittent execution relating to program or algorithm.				A11|A1103|A110302
		Program or Algorithm Execution Failure	A110303	Problem associated with the failure of a program or algorithm to execute. Sudden/unexpected interruption to a program's execution.				A11|A1103|A110303
	Computer Operating System Problem		A1104	Problem associated with software, firmware, and/or hardware elements that control the execution of computer programs and provides such services as computer resource allocation, job control, input/output control, and file management in a computer system.				A11|A1104
		Operating System Becomes Nonfunctional	A110401	Problem associated with malfunction of the computer operating system as opposed to an application software problem.				A11|A1104|A110401
		Operating System Version or Upgrade Problem	A110402	Problem associated with replacing an older operating system to an up-to-date operating system.				A11|A1104|A110402
	Computer System Security Problem		A1105	Problem associated with unauthorized access to or modification of a software system resulting in a loss of confidentiality, integrity, or availability of written program code, application software, or data or entire device.				A11|A1105
		Application Security Problem	A110501	Problem associated with the acquisition of computer programming codes that can replicate and spread from one computer system to another thereby leading to damaged software, hardware and data.				A11|A1105|A110501
		Unauthorized Access to Computer System	A110502	Problem associated with an access that was not permitted to the computer system that may lead to modification of program, corruption of data, or and break in network security. This concept is closely associated with computer integrity which is the degree to which a system or component prevents unauthorized access to, or modification of, computer programs or data.				A11|A1105|A110502
	Data Back-Up Problem		A1106	Problems relating to a system, component, file, procedure, or person available to replace or help restore a primary item in the event of a failure or externally caused disaster.				A11|A1106
		Failure to Back-Up	A110601	Problem associated with the inability to backup or to retrieve a backed up version (corrupted file) of device data or system files.				A11|A1106|A110601
		Failure to Convert to Back-Up	A110602	Problem associated with a failure to transition from a primary system, component, file, procedure to a backup in response to a failure in the primary item.				A11|A1106|A110602
	Data Problem		A1107	Event in which data (charting, orders, results) is not correctly stored, transferred, updated, or displayed.				A11|A1107
		Loss of Data	A110701	Event in which data is unintentionally permanently or temporarily lost, deleted, corrupted, or overwritten.				A11|A1107|A110701
		Patient Data Problem	A110702	Event in which data is accessed by the healthcare provider and either the wrong patient or the wrong data is retrieved despite correct inquiry procedures.				A11|A1107|A110702
	Date/Time-Related Software Problem		A1108	Problem associated with programming of calendar dates and/or time as a factor in the operation of the device.				A11|A1108
Connection Problem			A12	Problem associated with linking of the device and/or the functional units set up to provide means for a transfer of liquid, gas, electricity or data.				A12
	Blocked Connection		A1201	Problem associated with linking of the device whereby their functional units set up to provide means for a transfer of fluid, gas, or data is prevented or impeded.				A12|A1201
	Decoupling		A1202	Problem associated with the device being unassociated in such a way that fluid, gas, power or signal information may not be transferred from one to another.				A12|A1202
	Disconnection		A1203	Problem associated with the linking of the device having a sufficient open space to prevent gas, liquid or electrical current flow between connectors.				A12|A1203
	Failure to Disconnect		A1204	Problem associated with the linking of the device whereby termination of the transfer of liquid, gas, electricity, or information cannot be accomplished, or linking components do not come apart, or disconnect, when expected.				A12|A1204
	Loose or Intermittent Connection		A1205	Problem associated with the connection of the device being loose or intermittent.				A12|A1205
	Misconnection		A1206	Problem associated with the connection of the device being improper or not in accordance with device specification, requirements or intended uses.				A12|A1206
	Incomplete or Inadequate Connection		A1207	Problem associated with a partial linking of the device whereby device may appear to be connected however only a partial, intermittent or no transfer of liquid, gas, electricity, or information can be accomplished.				A12|A1207
	Fitting Problem		A1208	Problem associated with the connection of the device whereby channels, switching systems, and other functional units set up to provide means for a transfer of liquid, gas, electricity, or information do not match or fit.				A12|A1208
Communication or Transmission Problem			A13	Problem associated with the device sending or receiving signals or data. This includes transmission among internal components of the device to which the device is intended to communicate.				A13
	Failure to Read Input Signal		A1301	Problem associated with a failure of the device to read a signal for interpretation or measurement.				A13|A1301
	Failure to Transmit Record		A1302	Problem associated with a failure of the device to transmit a record for interpretation or measurement.				A13|A1302
	Intermittent Communication Failure		A1303	Inconsistent or lack of intended communication of data among internal components or with other external devices.				A13|A1303
	Telemetry Discrepancy		A1304	Problem associated with variability of the transmission of telemetry signals.				A13|A1304
	Wireless Communication Problem		A1305	Problems with the RF wireless technology characteristics and performance (e.g., frequency, output power, range, reception), wireless quality of service, wireless coexistence, security of wireless signals and data, and electromagnetic compatibility.				A13|A1305
Infusion or Flow Problem			A14	Problem associated with the device failing to deliver or draw liquids or gases as intended (e.g. delivering drugs at incorrect rate, Problems with drawing fluid from a system). This includes vacuum collection devices.		Modified (technical)	Technical change to the definition on 18 March 2020 to include vacuum collection devices.	A14
	Deflation Problem		A1401	Problem associated with the deflation of the device including unintended deflation or inadequate deflation.		Modified (technical)	Technical change to the definition on 18 March 2020 to accommodate additional level 3 terms.	A14|A1401
		Failure to Deflate	A140101	Problem associated with the inability to deflate a device which is intended to deflate.		New	Term was added on 18 March 2020 to differentiate between a failure to deflate and an unintended deflation.	A14|A1401|A140101
		Unintended Deflation	A140102	Problem associated with the unintended deflation of a device which was not designed to deflate or not at that point in time or not to that extent.		New	Term was added on 18 March 2020 to differentiate between a failure to deflate and an unintended deflation.	A14|A1401|A140102
	Excess Flow or Over-Infusion		A1402	Problem associated with a delivery overdose of therapeutic agents, such as drugs or fluids being delivered into a device or a patient.				A14|A1402
	Filling Problem		A1403	Problem associated with the method or amount of time associated with the delivery of a fluid. Time to delivery or amount of delivered entity may be affected.				A14|A1403
		Inability to Auto-Fill	A140301	Complete failure to fill as part of an automated process. For insufficient filling use "Short Fill". For excessive filling use "Overfill". For inconsistent filling use "Volume Accuracy Problem".				A14|A1403|A140301
		Overfill	A140302	Excessive filling of a device. For complete failure to fill use "inability to auto-fill". For insufficient filling use "Short Fill". For inconsistent filling use "Volume Accuracy Problem".				A14|A1403|A140302
		Short Fill	A140303	Insufficient filling of a device. For complete failure to fill use "inability to auto-fill". For insufficient filling use "short fill". For inconsistent filling use "Volume Accuracy Problem".				A14|A1403|A140303
		Volume Accuracy Problem	A140304	Inconsistent filling of a device. This describes a problem which is observed to vary between overfilling and under filling, and may be intermittent. Use "Overfill" or "Short Fill" if problem is consistent.				A14|A1403|A140304
	Filtration Problem		A1404	Problem associated with the process of passing a substance through a porous medium, e.g., a blood clot filter for the removal of suspended matter.				A14|A1404
		Inadequate Filtration Process	A140401	Problem associated with the filter failing to remove items or substances which should have been removed.				A14|A1404|A140401
		Inadequate Ultra Filtration	A140402	Problem associated with the transfer of fluid between the blood and dialysate through the dialysis membrane due to a pressure gradient (trans-membrane pressure) existing between the blood and dialysate compartments.				A14|A1404|A140402
	Improper Flow or Infusion		A1405	Problem associated with the regulation and delivery of therapeutic agents (e.g. air, gas, drugs or fluids into a device or a patient under positive pressure).				A14|A1405
		Backflow	A140501	Continuous flow of fluid (e.g. liquid, gas) against the intended flow direction.				A14|A1405|A140501
		Free or Unrestricted Flow	A140502	Problem associated with uncontrolled flow of infusion of air, gas or fluids.				A14|A1405|A140502
		Gradient Increase	A140503	Problem associated with the increased rate of change in temperature, pressure, or other variables as a function of distance, time, etc.				A14|A1405|A140503
		Inaccurate Delivery	A140504	Delivery at endpoint not as intended; either too low or too high.				A14|A1405|A140504
		Inaccurate Flow Rate	A140505	Problem associated with fluctuations in the flow volume delivered per time, even if end volume is correct, and delivered in the correct total time.				A14|A1405|A140505
		Intermittent Infusion	A140506	Problem associated with the infusion not being stead- characterised by intermittent stoppages to the flow.				A14|A1405|A140506
		Reflux within Device	A140507	Problem associated with partial backflow, compromising the device's flow output.				A14|A1405|A140507
		Restricted Flow rate	A140508	Problem associated with flow rate. Flow volume delivered over time is not reaching intended flow rate.				A14|A1405|A140508
		Tidal Volume Fluctuations	A140509	Problem associated with the amount of gas that is inspired and expired during one respiratory cycle.				A14|A1405|A140509
	Inflation Problem		A1406	Problem associated with the inability of the device to expand or enlarge with the intended inflation agent (e.g. saline or air).				A14|A1406
	Insufficient Flow or Under Infusion		A1407	Problem associated with an insufficient dose of therapeutic agents, e.g., drugs or fluids being delivered into a patient under positive pressure.				A14|A1407
	No Flow		A1408	Problem arising from the device failing to deliver the specified liquid or gas.				A14|A1408
		Failure to Deliver	A140801	Failure (=complete nonperformance) with regard to the intended function of delivery.				A14|A1408|A140801
		Failure to Infuse	A140802	Failure (=complete nonperformance) with regard to the intended function of infusion.				A14|A1408|A140802
		Inability to Irrigate	A140803	Failure (=complete nonperformance) with regard to the intended function of irrigation.				A14|A1408|A140803
	Obstruction of Flow		A1409	Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in restriction of flow (including blood clotting).		Modified (technical)	Technical change to the definition on 18 March 2020 to clarify that the term includes problems related to blood clotting.	A14|A1409
		Complete Blockage	A140901	Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in no flow.				A14|A1409|A140901
		Partial Blockage	A140902	Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in a reduction of the flow rate.				A14|A1409|A140902
	Difficult to Flush		A1410	The device is difficult to flush, possibly indicating an obstruction within device.				A14|A1410
	Pressure Problem		A1411	Problem associated with the application of a force either internal or external to device that compromises the flow of fluid or gas.				A14|A1411
		Decrease in Pressure	A141101	Unintended decrease in pressure, compromising the device's intended function.				A14|A1411|A141101
		Increase in Pressure	A141102	Unintended increase in pressure, compromising the device's intended function.				A14|A1411|A141102
		No Pressure	A141103	Unintended complete loss of pressure, compromising the device's intended function.				A14|A1411|A141103
	Pumping Problem		A1412	Problem associated with pump performance deviating from specifications in a way to compromise flow or infusion.				A14|A1412
		Decreased Pump Speed	A141201	Unintended decrease in pump speed and hence, probably, flow rate, compromising the intended function of the device.				A14|A1412|A141201
		Increased Pump Speed	A141202	Unintended increase in pump speed and hence, probably, flow rate, compromising the intended function of the device.				A14|A1412|A141202
		Failure to Pump	A141203	Problem associated with the device which fails to start pumping.				A14|A1412|A141203
		Pumping Stopped	A141204	Unexpected/unintended cessation of pump.				A14|A1412|A141204
	Suction Problem		A1413	Problem associated with suction equipment, which may be a manual, electrical, vacuum or pressure source operated to evacuate and remove undesired substances (air, gas, fluid, or particulates) via tubing and collection bag.				A14|A1413
		Decrease in Suction	A141301	Problem associated with the removal of fluid or gas from a body cavity or device due to decreased suction.		Modified (technical)	Technical change to the definition on 18 March 2020 to include removal of fluid or gas from a device.	A14|A1413|A141301
		Increase in Suction	A141302	Problem associated with the removal of fluid or gas from a body cavity or device due to increased suction.		Modified (technical)	Technical change to the definition on 18 March 2020 to include removal of fluid or gas from a device.	A14|A1413|A141302
		Suction Failure	A141303	Problem associated with the complete inability to provide suction.				A14|A1413|A141303
	Priming Problem		A1414	Problem associated with the preparation of the device to begin pumping.				A14|A1414
		Failure to Prime	A141401	Problem associated with the device failing to begin the priming process (i.e. the process of preparation of device for the delivery of fluids).				A14|A1414|A141401
		Incomplete or Inadequate Priming	A141402	Problem associated with not adequately preparing the device.				A14|A1414|A141402
	Air/Gas in Device		A1415	Problem associated with air in the device (e.g. air in line).		New	Term was added on 18 March 2020 as an additional level 2 term to "Infusion or Flow Problem".	A14|A1415
Activation, Positioning or Separation Problem			A15	Problem associated with any deviations from the documented specifications of the device that relate to the sequence of events for activation, positioning or separation of device. Note: Deployment is synonymous with activation.		Modified (editorial)	Editorial change to the term on 16 Feburary 2021.	A15
	Activation Problem		A1501	Problem associated with the activation of the device. Note: Activation includes expansion.		Modified (technical)	Technical change to the definition on 18 March 2020 to clarify that activation includes expansion.	A15|A1501
		Activation Failure	A150101	Problem associated with the device failing to be activated including partial activation.		Modified (editorial)	Editorial change to the definition on 18 March 2020.	A15|A1501|A150101
		Difficult or Delayed Activation	A150102	Problem associated with delayed or difficult activation of the device.				A15|A1501|A150102
		Premature Activation	A150103	Problem associated with early and unexpected activation of the device.				A15|A1501|A150103
		Self-Activation or Keying	A150104	Problem associated with the unintended activation of the device, or the device having been unexpectedly turned on during use.				A15|A1501|A150104
		Key or Button Unresponsive/not Working	A150105	Problem associated with a device not responding to key or button inputs.		New	Term was added on 18 March 2020 as an additional level 3 term to "Activation Problem".	A15|A1501|A150105
	Positioning Problem		A1502	Problem associated with the movement of the device to an intended location.				A15|A1502
		Positioning Failure	A150201	Problem associated with the inability of the device to be positioned in a specified location.				A15|A1502|A150201
		Malposition of Device	A150202	Problem associated with the device being positioned in a location other than intended or specified.				A15|A1502|A150202
		Difficult or Delayed Positioning	A150203	Problem associated with users experiencing difficulty or delay to position the device to a specified location.				A15|A1502|A150203
		Failure to Advance	A150204	Problem associated with failure to move the device to an intended location.				A15|A1502|A150204
		Difficult to Advance	A150205	Problem associated with difficulty moving the device to an intended location (e.g. difficulty in advancing guide wire).				A15|A1502|A150205
		Difficult to Insert	A150206	Problem associated with problems introducing or inserting the device, even if the user is operating the device in accordance with the instructions for use or labeling.				A15|A1502|A150206
		Difficult to Remove	A150207	Problem associated with the user experiencing difficulty when taking out the device, even if the user is operating device in accordance with the instructions for use or labeling.		Modified (editorial)	Editorial change to the definition on 18 March 2020.	A15|A1502|A150207
		Entrapment of Device	A150208	Problem associated with the device caught within patient vasculature, tissue, or other device.				A15|A1502|A150208
	Separation Problem		A1503	Problem associated with the detachment or separation of the device.				A15|A1503
		Separation Failure	A150301	Problem associated with the device or one of its components failing to detach or separate as intended.				A15|A1503|A150301
		Difficult or Delayed Separation	A150302	Problem associated with users experiencing difficulty or delay with detachment or separation of the device.				A15|A1503|A150302
		Premature Separation	A150303	Problem associated with an early and unexpected detachment or separation of the device from the system.				A15|A1503|A150303
Protective Measures Problem			A16	Problem associated with any deviations from the documented specifications of the device that relate to the implemented and inherited design features specific to devices used for reducing risks to patient or caregiver or maintaining risks within specified levels.				A16
	Device Alarm System		A1601	Problem associated with the alarm system of the device.				A16|A1601
		Alarm Not Visible	A160101	The device does not display an alarm message when required.				A16|A1601|A160101
		No Audible Alarm	A160102	The device fails to emit an audible alarm.				A16|A1601|A160102
		Low Audible Alarm	A160103	The audible device alarm cannot be heard clearly.				A16|A1601|A160103
		Delayed Alarm	A160104	The device alarm system operates with delay.				A16|A1601|A160104
		False Alarm	A160105	Problem associated with the device providing incorrect alarm warning or alert to user.				A16|A1601|A160105
		Defective Alarm	A160106	The device alarm does not operate as expected and/or in agreement with device's specifications.				A16|A1601|A160106
	Fail-Safe Problem		A1602	Problem associated with the feature that prevents the unsafe use of the device.				A16|A1602
		Fail-Safe Did Not Operate	A160201	Problem associated with the device fail-safe mechanism, which did not function or function in a non effective way, compromising safe use of the device.				A16|A1602|A160201
		No Fail-Safe Mechanism	A160202	The device does not have a fail-safe mechanism, although such mechanism would be required for its appropriate and/or safe functioning.				A16|A1602|A160202
	Failure of Device to Self-Test		A1603	Problem associated with the device failing to perform an internal self-diagnostic process to ensure normal operation during or prior to use.				A16|A1603
	Failure to Auto Stop		A1604	Problem associated with the inability of device to turn itself off when the device is not in an operable condition.				A16|A1604
	Reset Problem		A1605	Problem associated with setting a variable, register, or other storage location back to a prescribed state.				A16|A1605
		Failure to Reset	A160501	Problem associated with the device failing to set a variable, register, or other storage location back to a prescribed state.				A16|A1605|A160501
		Failure to Zero	A160502	Problem associated with the device failing to set a variable, register, or other storage location back to zero.				A16|A1605|A160502
		Inappropriate or Unexpected Reset	A160503	Problem associated with the device setting a variable, register, or other storage location to an inappropriate or unexpected state.				A16|A1605|A160503
	Premature Indicator Activation		A1606	Problems with the activation of a protective measure indicator earlier than expected.				A16|A1606
		Premature Elective Replacement Indicator	A160601	Problems with the early or unexpected activation of the elective replacement indicator.				A16|A1606|A160601
		Premature End-of-Life Indicator	A160602	Problem with the early or unexpected activation of the end-of-life indicator.				A16|A1606|A160602
	Shielding Failure		A1607	Problem associated with the device inability to act as a barrier for absorption of radiation energy in X-rays, gamma rays, etc.				A16|A1607
Compatibility Problem			A17	Problem associated with compatibility between device, patients or substances (medication, body fluid, etc.)				A17
	Component or Accessory Incompatibility		A1701	Problem associated with the incompatibility of any device while being operated in the same use environment thereby leading to a dysfunction between the devices.				A17|A1701
		Accessory Incompatible	A170101	An accessory required for the intended purpose of the device appears incompatible with device, thus compromising the intended function of the device.				A17|A1701|A170101
		Component Incompatible	A170102	A component required for the proper functioning of the device is not compatible with other components or subassemblies of the device, thus compromising the intended function of the device.				A17|A1701|A170102
	Device-Device Incompatibility		A1702	Problem associated with the incompatibility of two or more devices while being operated in the same use environment thereby leading to a dysfunction of more than one device.				A17|A1702
	Measurement System Incompatibility		A1703	Problem associated with the incompatibility of the measurement systems between and/or within device systems that are inherent to the individual device thereby leading to miscalculated or mismatched measurements from those devices, e.g., international metric system versus U.S. measurement system.				A17|A1703
	Unintended Compatibility		A1704	Problem associated with the ability of two or more devices which are intended to be incompatible but are able to work or fit together.				A17|A1704
Contamination /Decontamination Problem			A18	Problem associated with the presence of any unexpected foreign substance found in the device, on its surface or in the package materials, which may affect performance or intended use of the device, or problem that compromise effective decontamination of the device.				A18
	Contamination		A1801	Problem associated with the undesired introduction of impurities either chemical or microbiological in nature, or of foreign matter into or onto the device.		Modified (technical)	Technical change to the term on 18 March 2020 to separate when/where contamination occurred from type of contamination ("During Use" was removed). NOTE: This change affects all lower level terms (A180101 - A180105) which are also no longer limited to contamination during use.	A18|A1801
		Biofilm coating in Device	A180101	Problem associated with the undesired introduction of a biofilm coating into or onto the device.				A18|A1801|A180101
		Contamination of Device Ingredient or Reagent	A180102	Problem associated with the undesired introduction of impurities either chemical or microbiological in nature, or of foreign matter into or onto the device ingredient or reagent.				A18|A1801|A180102
		Device Contamination with Body Fluid	A180103	Problem associated with the undesired presence of body fluid in/on the device, which are not part of the documented device specifications and requirements.				A18|A1801|A180103
		Device Contamination with Chemical or Other Material	A180104	Problem associated with contamination of the device with chemical substance or other non biologic material.				A18|A1801|A180104
		Microbial Contamination of Device	A180105	Problem associated with undesired microbial contamination of the device.				A18|A1801|A180105
	Device Contaminated During Manufacture or Shipping		A1802	Problem associated with the presence of any unexpected foreign substance found on the surface or in the package materials, which may affect optimal performance for its intended use, and which occurred during manufacture or shipping.		Modified (technical)	Technical change to the definition on 18 March 2020 to clarify that contamination occurred during manufacture or shipping.	A18|A1802
	Device Reprocessing Problem		A1803	Problem associated with a failure during any step of reprocessing process (cleaning, disinfection, packaging, labeling, sterilization) of a used or opened from its original packaging, but unused, device.				A18|A1803
		Failure to Clean Adequately	A180301	Problem associated with the failure of the device or operator to remove any visible soil, foreign material or organism deposits on the external surfaces, crevices, and joints of the device.				A18|A1803|A180301
		Failure to Disinfect	A180302	Failure to properly disinfect the device when reprocessing it.				A18|A1803|A180302
		Flushing Problem	A180303	Flushing process was not executed properly.				A18|A1803|A180303
		Problem with Removal of Enzymatic Cleaner	A180304	Enzymatic cleaner was not removed properly.				A18|A1803|A180304
		Problem with Sterilization	A180305	Device was not sterilized properly during reprocessing.				A18|A1803|A180305
		Residue After Decontamination	A180306	Problem associated with the decontamination process not adequately removing unwanted visible soil, foreign material, or organism deposits.				A18|A1803|A180306
	Device Contaminated at the User Facility		A1804	Problem associated with the presence of any unexpected foreign substance found on the surface or in the package materials, which may affect optimal performance for its intended use, and which occurred at the user facility.		New	Term was added on 18 March 2020 for coding of contamination at the user facility which was originally included in "Contamination During Use" (A1801).	A18|A1804
Environmental Compatibility Problem			A19	Problem associated with the surrounding conditions in which the device is being used such as temperature, noise, lighting, ventilation, or other external factors such as power supply.				A19
	Ambient Noise Problem		A1901	Problem associated with any undesired acoustic energy or vibration that tends to interfere with the operation of the device.				A19|A1901
	Ambient Temperature Problem		A1902	Problem associated with compromised device performance at the ambient temperature or the storage at an inappropriate ambient temperature.				A19|A1902
	Fumes or Vapors		A1903	Problem associated with the visibility, odor, or toxicity of an ambient vapor or gas.				A19|A1903
	Fungus in Device Environment		A1904	Problem associated with the visibility of molds, mildews, yeasts, and/or mushrooms in the immediate environment in which the device is being used.				A19|A1904
	Moisture or Humidity Problem		A1905	Problem associated with an unsatisfactory humidity level in the storage or use environment which affects the device performance.				A19|A1905
		Moisture Damage	A190501	Problem associated with damage inflicted upon the device from water vapor or water in the immediate environment in which the device is being used.				A19|A1905|A190501
	Ventilation Problem in Device Environment		A1906	Problem associated with the circulation of fresh air in the immediate atmosphere in which the device is being used.				A19|A1906
		Fogging	A190601	Problem associated with the visibility of water vapor in the immediate atmosphere in which the device is being used.				A19|A1906|A190601
	Device Unsafe to Use in Environment		A1907	Problem associated with environmental condition that results in the unsafe use of the device. (e.g. electromagnetic fields, noise, vibration, microbiological contamination etc.)				A19|A1907
	Environmental Particulates		A1908	Problem associated with fine solids or liquid particles such as dust, smoke, fume, and/or mist suspended in the immediate atmosphere in which the device is being used.				A19|A1908
	Medical Gas Supply Problem		A1909	Problem associated with the facility-supplied medical gases such as medical air, oxygen, nitrous oxide, and nitrogen.				A19|A1909
	Electrical Power Problem		A1910	Problem associated with the quality of the facility-supplied power.				A19|A1910
		Emergency Power Failure	A191001	Problem associated with the failure of the facility's emergency power backup system(s) including generators and/or interruptible power systems (UPS).				A19|A1910|A191001
		Loss of Power	A191002	Problem associated with the failure of primary power supplied by the facility.				A19|A1910|A191002
		Power Conditioning Problem	A191003	Problem associated with a momentary overpower/over voltage from the utility and electrical systems of user facilities; - Problem associated with inadequate power conditioning such as the presence of fluctuation, surges, spikes, dropouts, noise and other such undesirable transients.				A19|A1910|A191003
Installation-Related Problem			A20	Problem associated with unsatisfactory installation, configuration, and/or setup of a specific device.				A20
	Misassembled During Installation		A2001	Problem associated with the use of the device characterized by incorrect assembly of device components, parts or constituents.				A20|A2001
Labelling, Instructions for Use or Training Problem			A21	Problem associated with device markings/labelling, instructions for use, training and maintenance documentation or guidelines.				A21
	Device Markings/Labelling Problem		A2101	Problem associated with the written, printed or graphic material accompanying or affixed to the device or any of its packaging. This includes verbal instructions relating to identification, technical description, and usage provided by the device manufacturers. Problems can include but are not limited to this material being unclear, missing, worn out, incorrect or inaccurate.				A21|A2101
		Expiration Date Error	A210101	Problem associated with errors in identification of expiration date.				A21|A2101|A210101
		Illegible Information	A210102	Problem associated with information unable to be read or deciphered.				A21|A2101|A210102
		Inaccurate Information	A210103	Problem associated with imprecise, inexact information.				A21|A2101|A210103
		Unclear Information	A210104	Problem associated with ambiguous, confused information.				A21|A2101|A210104
		Missing Information	A210105	Absence of information e.g. labeling, instruction for use.				A21|A2101|A210105
	Lack of Maintenance Documentation or Guidelines		A2102	Problem associated with user facility not receiving adequate service documentation, guidelines, or recommendations to perform preventative and corrective maintenance and performance assurance checks.				A21|A2102
	Inadequate Instructions for Healthcare Professional		A2103	Problem associated with inaccuracies in any written, printed, or graphic matter that is affixed to the device or its packaging with any matter that accompanies the device including verbal instructions related to identification, technical description and use of device provided by the device manufacturers that is intended for healthcare professionals.				A21|A2103
	Inadequate Instructions for Non-Healthcare Professional		A2104	Problem associated with users being unclear and not able to follow any written, printed, or graphic matter that is affixed to device or its packaging with any matter that accompanies the device including verbal instructions related to identification, technical description and use of the device provided by the device manufactures that vary from the standard of medical care in a given environment.				A21|A2104
	Inadequate or Insufficient Training		A2105	Problem associated with facility not providing satisfactory initial and/or periodic user training covering operation of the device.				A21|A2105
Human-Device Interface Problem			A22	Problem associated with an act or omission of an act that has a different result than that intended by the manufacturer or expected by the operator.				A22
	Device Difficult to Setup or Prepare		A2201	Problem associated with the use of the device in terms of user experiencing difficulty in preparing device for use, even if the operation is being performed according to labeled instructions for use.				A22|A2201
	Device Difficult to Program or Calibrate		A2202	The device is difficult to program, calibrate or set to desired state, even by appropriately trained user/operator.				A22|A2202
	Device Difficult to Maintain		A2203	Problem associated with the user's ability to service the device according to the manufacturer specifications relating to the device routine maintenance, i.e., periodic inspection, failure detection, repair, and care of the device to sustain or restore acceptable operating conditions.				A22|A2203
	Inadequate User Interface		A2204	Problem associated with the means by which the operator and the equipment communicate or interact.				A22|A2204
Use of Device Problem			A23	Problem associated with failure to process, service, or operate the device according to the manufacturer's recommendations or recognized best practices.				A23
	Device Handling Problem		A2301	Handling of the device not in accordance with specification, prior to use on the patient.				A23|A2301
	Use of Incorrect Control Settings		A2302	Problem associated with the use of the device in terms of inappropriate and false control setting for the device's specified operation and/or intended use.				A23|A2302
	Improper or Incorrect Procedure or Method		A2303	Problem associated with the use of the device in terms of nonconforming to that device's intended use, specifications, procedure and process or service instructions and information provided by the device manufacturers.				A23|A2303
	Off-Label Use		A2304	Problem associated with the device which has been used for an unapproved indication or for an unapproved intended use.				A23|A2304
	Misassembled		A2305	Problem associated with incorrect assembly of the device or constituents after being put into use.				A23|A2305
		Misassembly by Users	A230501	Problem associated with incorrect assembly of the device or constituents by the users.				A23|A2305|A230501
		Misassembly During Maintenance/Repair	A230502	Problem associated with incorrect assembly of the device or constituents during maintenance or repair.				A23|A2305|A230502
Adverse Event Without Identified Device or Use Problem			A24	An adverse event (e.g. patient harm) appears to have occurred, but there does not appear to have been a problem with the device or the way it was used.				A24
No Apparent Adverse Event			A25	A report has been received but the description provided does not appear to relate to an adverse event. This code allows a report to be recorded for administration purposes, even if it doesn't meet the requirements for adverse event reporting.				A25
Insufficient Information			A26	An adverse event appears to have occurred but there is not yet enough information available to classify the device problem.				A26
Appropriate Term/Code Not Available			A27	The device problem is not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table.				A27